## Slide 1
Respiratory Oscillometry
Chiara Veneroni
Department of Electronics, Informatics and Biomedical Engineering
(chiara.veneroni@polimi.it)
Notes: Come si esegue la oscillometria step-by-step - 30min
La oscillometria nel paziente stabile ed acuto - 20 min
Esercitazioni oscillometria e interpretazione risultati - 1h??

## Slide 2
Outline
1. What is Oscillometry
Understanding the fundamental principles
Interpretation of results
How to perform an oscillometry test
Oscillometry in research and clinics
2. Main clinical applications of oscillometry
Detection of airway obstruction
Bronchial reversibility and Bronchial provocation test
Detection of Tidal Expiratory Flow limitation –EFLt
Other applications
Come si esegue la oscillometria step-by-step
La oscillometria nel paziente stabile ed acuto

## Slide 3
Section 1
What is Oscillometry?
Understanding the fundamental principles

## Slide 4: Measuring mechanical properties of the respiratory system
respiratory muscles
pressure
Flow
Respiratory mechanics: the determination of the relationships between pressure and flow in the respiratory system
Pressure acting on the lung during spontaneous breathing:
Pressure generated by muscles
Plung = Pmus
??? ???

## Slide 5: Forced Oscillation Technique or Oscillometry
respiratory muscles
pressure
Pressure
Flow
Oscillometry allows the determination of the lung mechanical properties by looking at the flow response to small amplitude, high frequency pressure oscillations
DuBois AB, Brody AW, Lewis DH, Burgess BF. Oscillation mechanics of lungs and chest in man. J Appl Physiol 1956; 8: 587-94
DuBois AB, Brody AW, Lewis DH, Burgess BF. Oscillation mechanics of lungs and chest in man. J Appl Physiol 1956; 8: 587-94

## Slide 6: Forced Oscillation Technique or Oscillometry
respiratory muscles
pressure
Pressure
Flow
Oscillometry allows the determination of the lung mechanical properties by looking at the flow response to small amplitude, high frequency pressure oscillations
DuBois AB, Brody AW, Lewis DH, Burgess BF. Oscillation mechanics of lungs and chest in man. J Appl Physiol 1956; 8: 587-94
DuBois AB, Brody AW, Lewis DH, Burgess BF. Oscillation mechanics of lungs and chest in man. J Appl Physiol 1956; 8: 587-94

## Slide 7: Forced Oscillation Technique or Oscillometry
respiratory muscles
pressure
Pressure
Flow
Oscillometry allows the determination of the lung mechanical properties by looking at the flow response to small amplitude, high frequency pressure oscillations
DuBois AB, Brody AW, Lewis DH, Burgess BF. Oscillation mechanics of lungs and chest in man. J Appl Physiol 1956; 8: 587-94
DuBois AB, Brody AW, Lewis DH, Burgess BF. Oscillation mechanics of lungs and chest in man. J Appl Physiol 1956; 8: 587-94

## Slide 8: Oscillometry: principles
respiratory muscles
pressure
Pressure
Flow
DuBois AB, Brody AW, Lewis DH, Burgess BF. Oscillation mechanics of lungs and chest in man. J Appl Physiol 1956; 8: 587-94
DuBois AB, Brody AW, Lewis DH, Burgess BF. Oscillation mechanics of lungs and chest in man. J Appl Physiol 1956; 8: 587-94
Pressure acting on the lung during spontaneous breathing:
Pressure generated by muscles AND device
Plung = Pmus + Pdevice
Oscillometry allows the determination of the lung mechanical properties by looking at the flow response to small amplitude, high frequency pressure oscillations
??? ???

## Slide 9
Forced Oscillation Technique or Oscillometry
If the oscillations are faster than muscles (=higher frequency) it is possible to split the two components and consider the flow resulting from the oscillatory pressure only.
+
Hz:
the number of times pressure oscillates in a second.
Oscillometry: principles

## Slide 10: OPTIMAL STIMULATION FREQUENCY: SPEED VS DEPTH
The proportion of the oscillatory pressure applied at the airway opening that is transmitted to the alveolus is higher at low frequencies than at higher frequencies.
Airway opening
6 Hz
12 Hz
Large airways
Alveoli

## Slide 11: OPTIMAL STIMULATION FREQUENCY: SPEED VS DEPTH
The optimal stimulation frequency is a trade-off between being faster than the breathing but not too fast, in order to reach the deepest parts of the lung.
For spontaneous breathing:
Adults: Frequency ≥ 5 Hz
Preschool children: Frequency ≥ 8 Hz
Newborns: Frequency ≥ 10 Hz
Which forcing frequencies should we use?

## Slide 12
What is Oscillometry?
Understanding the fundamental principles
Interpretation of results

## Slide 13
The oscillatory pressure moves air in and out of the lung, generating an airflow through the airways that is measured and used to calculate the respiratory impedance (Zrs).
Respiratory Impedance
(Zrs)
Resistance
Rrs
Reactance
Xrs
Impedance (Zrs) tells us how easy or difficult it is for air to flow in and out of the lungs.
Impedance has 2 components:
Resistance (Rrs) and Reactance (Xrs)
for each frequency (f) used for testing.
Respiratory impedance: INTERPRETATION

## Slide 14: RESISTANCE (Rrs): CLINICAL INTERPRETATION
Resistance (Rrs) accounts for obstruction of the central (larger) airways
Rrs becomes bigger in obstructive diseases
Resistance (Rrs) provides information related to the caliber of the larger airways, measured at tidal volume.
Rrs increases in situations such as:
  during bronchoconstriction,
  the presence of excessive mucous,
  airway inflammation,
  other causes of airway narrowing or obstruction1
1Kaminsky, D. A., et al. (2022). Clinical significance and applications of oscillometry. European Respiratory Review, 31(163), 1–19.
2Mead, J. (1970). The Lung’s Quiet Zone. New England Journal of Medicine, 282(23), 1318–1319.

## Slide 15: Small airways and airway resistance
F. Thien. Measuring and imaging small airways dysfunction in asthma. http://dx.doi.org/10.5415/apallergy.2013.3.4.224
Small airways

## Slide 16
Rrs
Example: Rrs in a normal lung
Oscillometry: the oscillogram

## Slide 17
Example: Rrs in a patient with central airway obstruction
Oscillometry: the oscillogram
Rrs

## Slide 18
Example: Rrs in a patient with peripheral AND/OR heterogeneous airway obstruction
Oscillometry: the oscillogram

## Slide 19
Example: Rrs in a patient with peripheral AND/OR heterogeneous airway obstruction
Oscillometry: the oscillogram
Rrs increases more at low frequencies:
«frequency dependence of Rrs»
Warning: frequency dependence is a sensitive but not specific marker of peripheral airway obstruction! Central heterogeneous obstruction and upper airway shunt also increase frequency dependence.

## Slide 20: REACTANCE(Xrs): CLINICAL MEANINGS
At low stimulating frequencies (f < Fres), Xrs is a negative number and dominated by the dynamic compliance of the respiratory system
At higher stimulating frequencies (f > Fres), Xrs is a positive number and dominated by inertial properties of respiratory system
«low frequency Xrs:
important for clinical
evaluation

## Slide 21: REACTANCE(Xrs): CLINICAL MEANINGS
SIMPLIFYING THE CONCEPT:
Xrs enables to determine how effectively the lung is being ventilated in its distal areas or how well air reaches the peripheral areas4.
Xrs falls below predicted values at low oscillating frequencies in the following conditions:
  peripheral obstruction1
  presence of tidal expiratory flow limitation (during expiration only)2
  alveolar gas trapping and/or closing of alveolar units3,4
Reactance reflects how easily the air reaches the lung periphery
Xrs becomes more negative in the presence of small airway disease or reduced lung/chest wall compliance
1Downie, S. R. et al(2013). Journal of Applied Physiology (Bethesda, Md. : 1985), 114(6), 770–777.
2Dellacà, R. L. et al. (2004). The European Respiratory Journal, 23(2), 232–240.
3Dellacà, R. L. et al (2009). Intensive Care Medicine, 35(12), 2164–2172.
4Milne, S. et al (2019). Journal of Applied Physiology (Bethesda, Md. : 1985), 126(5),1223–1231
At low FOT frequencies (5 Hz – 8 Hz), Xrs is related to the amount of alveoli in communication with the airway opening.

## Slide 22: REACTANCE(Xrs): CLINICAL MEANINGS
Low-frequency Xrs (i.e., 5-8 Hz) becomes ABNORMAL when the measurable part of the respiratory system, including the lungs and chest wall, becomes “stiffer” (i.e., dynamic compliance is reduced).
When there is an increase in the stiffness of the chest wall, which can occur in disorders such as osteogenesis imperfecta or amyotrophic lateral sclerosis (ALS).
When there is an increase in the stiffness of the lung parenchyma, as seen in Interstitial lung diseases
When oscillations cannot reach the periphery, such as in cases of small airway obstruction/collapse, or alveolar derecruitment.
Factors Contributing to a Reduction in Dynamic Compliance:
PERIPHERAL ʺROADBLOCKSʺ
STIFFER WALLS
STIFFER TISSUE

## Slide 23
Example: Rrs in a patient with lung volume de-recruitment (gas trapping, atelectasis, etc)
Oscillometry: the oscillogram
Rrs
Xrs

## Slide 24
Example: Rrs and Xrs in a patient with peripheral airway obstruction (eventually associated with lung volume de-recruitment and heterogeneous obstruction)
Oscillometry: the oscillogram

## Slide 25: REACTANCE(Xrs): CLINICAL INTERPRETATION
SIMPLIFYING THE CONCEPT:
1Downie, S. R. et al(2013). Journal of Applied Physiology (Bethesda, Md. : 1985), 114(6), 770–777.
2Dellacà, R. L. et al. (2004). The European Respiratory Journal, 23(2), 232–240.
3Dellacà, R. L. et al (2009). Intensive Care Medicine, 35(12), 2164–2172.
4Milne, S. et al (2019). Journal of Applied Physiology (Bethesda, Md. : 1985), 126(5),1223–1231

## Slide 26: OSCILLOMETRY: KEY PARAMETERS
R5-19
R5
Ax
Fres
X5
  Rrs and Xrs are usually expressed as cmH2O/(L/s) or hPa/(L/s)
is always a positive number!
is «flat» on healthy subjects
shows a frequency dependence in disease
Rrs:
Commonly used parameters:
R5: Rrs at 5 Hz (or R8, 8 Hz, in preschooler)
R5-19: difference between Rrs at 5 Hz and 19 Hz, named also ‘frequency dependance of Rrs’
usually negative at low frequencies
positive at higher frequencies
Xrs:
  Commonly used parameters:
  X5: Xrs at 5 Hz (or X8, 8 Hz, in preschooler)
  Fres: frequency where Xrs=0
  Ax: area of Xrs between 5 Hz and Fres
PRACTICAL HINT: USE only a single Xrs parameter to describe lung periphery
X5, Ax and Fres provide equivalent information related to peripheral airway disease1
1Veneroni C et al. Diagnostic Potential of Oscillometry: A Population-based Approach. Am J Respir Crit Care Med. 2024;209(4):444-453.

## Slide 27
insp
Flow (ml/s)
Time (s)
With FOT
oscillatory signal
Oscillometry: within breath changes

## Slide 28: Airway dynamics during inspiration and expiration
Expiration
Inspiration
During inspiration, respiratory muscles expand the lung to reduce the alveolar pressure and initiate a flow of fresh air from the airway opening to the alveoli. During this phase, airway transmural pressure dilates the airways.
During expiration, the elastic recoil of the lung and rib cage (and, during exercise, the expiratory muscular activity) increases alveolar pressure to move the air from the alveoli to the airway opening. During this phase, airway transmural pressure compresses the airways (“dynamic airway compression”).
Notes: Cambiano I volume… cambiano le prop meccanchie
Ma non solo I volume sono anche diverse le forze che agiscono sulle vie aeree durante insp/esp

## Slide 29: Tidal EFL (EFLt): physiology
TIDAL EFL: the expiratory flows are so reduced that the lung cannot empty before the beginning of a new breath
“Expiratory flow limitation (EFL) is the pathophysiological hallmark of chronic obstructive pulmonary disease (COPD) and arises because of the dual effects of permanent parenchymal destruction (emphysema) and airway dysfunction, which in turn reflect the effects of small airway inflammation (mucosal oedema, airway remodelling and mucous impaction) and possibly increased cholinergic airway smooth muscle tone”
Physiology and consequences of lung hyperinflation in COPD.
D. E. O'Donnell, P. Laveneziana
European Respiratory Review 2006 15: 61-67;
Normal
COPD
Notes: 1- enfisema,.. Riduzione delle forze di richiamo eleastico
2- infiammazione, edema, smoth muscle.. La forza non si trasmette sul diametro interno la perdo sulle pareti

## Slide 30: Within breath Oscillometry: KEY PARAMETERS
In within-breath oscillometry, Rrs and Xrs are studied in their temporal changes during the breathing cycle
Common parameters used in within-breath oscillometry (calculated at the lowest available frequency):
Rinsp: average Rrs during inspiration
Xinsp: average Xrs during inspiration
Rexp: average Rrs during expiration
Xexp: average Xrs during expiration
ΔXrs: Xinsp-Xexp
Inspiration
Rinsp
Rexp
ΔXrs
Expiration
Respiratory diseases produce specific within-breath changes in the mechanical properties of the lung.

## Slide 31: Within breath Oscillometry: ADVANTAGES
Whole-breath analyses show increased R5 and decreased X5 in the asthma, COPD and ILD groups.
Changes in within breath measurement results, particularly X5 can show specific patterns in asthma, COPD, and ILD.
Sugiyama A, Hattori N, Haruta Y, et al. Characteristics of inspiratory and expiratory reactance in interstitial lung disease. Respir Med. 2013;107(6):875-882.
KEY MESSAGE
Within-breath oscillometry can measure alterations that are more evident during a specific breathing phase.

## Slide 32
What is Oscillometry?
Understanding the fundamental principles
Interpretation of results
How to perform an oscillometry test

## Slide 34: Oscillometry: measurement protocol
> 3 artifacts free breaths
rest
> 3 artifacts free breaths
rest
> 3 artifacts free breaths
60-90 s
30 s
60-90 s
30 s
60-90 s
Measurement 1
Measurement 2
Measurement 3
What is an oscillometry evaluation?
average of at least 2 measurements with within-session coefficient of variability (CoV) of Rtot <10%(adults) or < 15% (children)
Source: Technical standards for respiratory oscillometry. Eur Respir J. 2020;
Average duration of a spirometry test: 30 minutes
(Am J Respir Crit Care Med, Vol. 189, P17-P18, 2014)
For maximum sensitivity, it's essential to conduct oscillometry testing BEFORE other Pulmonary Function Tests (PFTs) that necessitate the performance of Deep Inhalation.
Notes: Ancora meglio su numero di respire interi

## Slide 35
Perform the test with the patient correctly seated, with an upright posture with a slight “chin-up” position.
Instruct the patient to grip the mouthpiece with teeth and lips firmly to avoid leaks and to keep the tongue relaxed and below the mouthpiece, avoiding blockage.
Support the patient cheeks, and the floor of the mouth with palm and fingers or demonstrate to the patient how to do it by himself
Running the test
Anticipate they may experience a “gentle vibration” or “fluttering” during the test
Anticipate that the test will last approximately the time required to record 5 to 10 quiet breaths and that it will be repeated 3 times
Encourage the patient to stay relaxed and breathe as normal

## Slide 36: Evaluating possible measurement artifacts
Only artifact-free portions of the measurements should be used for the processing of final results.
- Marchal F, et al.. Filtering artifacts in measurements of forced oscillation respiratory impedance in young children. Physiol Meas. 2004;25(5):1153-1166.
- King GG, et al.. Technical standards for respiratory oscillometry. Eur Respir J. 2020;55(2):1900753.
Aitken CR, et al. Reproducibility and responsiveness of airway impedance measures derived from the forced oscillation technique across different operating lung volumes. Respir Physiol Neurobiol. 2024;320(September 2023):104200.
Oscillometry may be affected by some measurement artifacts such as:
RELATED TO UPPER AIRWAYS
Missing cheek support
coughing
glottis closure
swallowing,
tongue obstruction of the filter (especially in children)
POSTURE RELATED
Leaks around the filter
Leaks around the mouthpiece/forget to use the noseclip
Incorrect measurement position
BREATHING PATTERN RELATED
Hyperventilation
High tidal volumes (i.e. > 2x resting tidal volume)

## Slide 37
What is Oscillometry?
Understanding the fundamental principles
Interpretation of results
How to perform an oscillometry test
Oscillometry in research and clinics

## Slide 38
Prof. Arthur Dubois introduced FOT.
(Journal of Applied Physiology 1956)
Development of simpler and commercially available FOT devices
ERS task force guideline document
Eur Respir J 2003; 22: 1026–1041
ERS technical standard
Eur Respir J. 2020;55(2):1900753.
ATS/ERS statement document (preschooler)
Am J Respir Crit Care Med. 2007;175(12):1304-1345.
Oscillometry: development
New publications that demonstrate the clinical utility of oscillometry
Pubblication of technical standard documents

## Slide 39: Advantages of oscillometry in clinical practice
Faster testing
Oscillometry testing requires 3 measurements of 5-10 breaths each for a total of about 5 minutes.
Less effort required by patients
Oscillometry requires less effort from the patient, making it more suitable for children and elderly patients.
Very sensitive to changes in lung function
Oscillometry is more sensitive to subtle changes in lung function, making it better for detecting early signs of respiratory disease.
Comprehensive
Oscillometry provides a comprehensive assessment of lung function, including small airway function, that spirometry may miss.

## Slide 40: Conclusions
“.. oscillometry has gained considerable attention in the last two decades with the introduction of commercially available instruments, the publication of updated technical standards, and renewed enthusiasm for an old technique. ”
“… validating oscillometry as an important complementary, and advantageous, tool to spirometry in the assessment of respiratory symptoms and diagnosis of lung disease.”

## Slide 41
Outline
1. What is Oscillometry
Understanding the fundamental principles
Interpretation of results
How to perform an oscillometry test
Oscillometry in research and clinics
2. Main clinical applications of oscillometry
Detection of airway obstruction
Bronchial reversibility and Bronchial provocation test
Detection of Tidal Expiratory Flow limitation –EFLt
Other applications
Come si esegue la oscillometria step-by-step
La oscillometria nel paziente stabile ed acuto

## Slide 42
Outline
2. Main clinical applications of oscillometry
Detection of airway obstruction
Bronchial reversibility and Bronchial provocation test
Detection of Tidal Expiratory Flow limitation (EFLt)
Other applications

## Slide 43: Detection of central and peripheral obstruction
A number of normative ranges have been published for Rrs and Xrs parameters
GLI is currently working to the difinition of global reference equation for oscillometry
Eur Respir J. 2020;55(2):1900753.
Notes: Problema di quale scegliere e ai 18 anni
.. Sulla popolazione poco impatto.. Sul singolo sì..
Lettera ricevuta..
Non possiamo trascurare le differenze anatomiche.. Xrs volume, Rrs diametric… se il polmone è più grande sono più belle.

## Slide 44: Clinical interpretation of an oscillometry test
↔ : Within normal range
↗: above ULN
↘: below LLN
R5 | X5 | Interpretation:
↔ | ↔ | Normal respiratory mechanics
↗ | ↔ | Central (upper?) airway obstruction
↔ | ↘ | Peripheral airway obstruction and/or lung volume de-recrutment / gas trapping
↗ | ↘ | Heterogeneous (mixed) obstruction

## Slide 45: Clinical interpretation of an oscillometry test
↔ : Within normal range
↗: above ULN
↘: below LLN
R5 | X5 | Ax | Fres | Interpretation:
↔ | ↔ | ↔ | ↔ | Normal respiratory mechanics
↗ | ↔ | ↔ | ↔ | Central (upper?) airway obstruction
↔ | ↘ | ↗ | ↗ | Peripheral airway obstruction and/or lung volume de-recrutment / gas trapping
↗ | ↘ | ↗ | ↗ | Heterogeneous (mixed) obstruction
Correlated variables
References: - Veneroni et al - Diagnostic potential of forced oscillometry technique : a population-based approach American Journal of Respiratory and Critical Care Medicine, Volume 209 Number 4 | February 15 2024 - Neeraj Gupta - Can oscillometry be a substitute to spirometry in the indigent? - Lung India 2022;39:74-6.

## Slide 46
General unselected population 18-90 years: 7.560 subjects
Abnormalities in oscillometry parameters are significantly related to the presence of respiratory symptoms and previous diagnosis of respiratory disease.
20% the adult population have abnormal oscillometry parameters, 13% of the adult population have abnormal spirometry
Abnormalities in oscillometry parameters are related to symptoms and previous respiratory diagnosis even in subjects with normal spirometry
Am J Respir Crit Care Med. 2024;209(4):444-453.

## Slide 47
FEV1/FVC >LLN or FVC >LLN
COMPLETE STUDY POPULATION

## Slide 48: Oscillometry and disease progression
1 Di Mango et al. (2006). Changes in respiratory mechanics with increasing degrees of airway obstruction in COPD: detection by forced oscillation technique. Respiratory Medicine, 100(3), 399–410.
KEY MESSAGE:
Rrs and Xrs can be used as indicators of changes in the level of obstructions in both the larger and smaller airways, which occur as the disease progresses.1
As the disease worsens:
Rrs at 5 Hz become more POSITIVE
Xrs at 5 hz become more NEGATIVE
control
severe obstruction
control
severe obstruction
OSCILLOGRAM:
Rrs and Xrs values measured at different oscillatory “frequencies”
Notes: Gruppi molto ben alliati per età, altezza e peso

## Slide 49
Main clinical applications of oscillometry
Detection of airway obstruction
Bronchial reversibility and Bronchial provocation test
Detection of Tidal Expiratory Flow limitation (EFLt)
Other applications
Clinical applications of oscillometry

## Slide 50: Bronchodilator reversibility testing
What is the purpose of a bronchodilator reversibility (BDR) test?
When an airway obstruction is detected, a BDR test aids physicians in understanding how obstructed airways react to the administration of a bronchodilator drug.
It is recommended1,2 to support diagnosis of Asthma and Chronic Obstructive Pulmonary Disease (COPD) and for evaluating the response to medication.
*Historically perfomed using spirometry
By BruceBlaus - Own work, CC BY-SA 4.0, https://commons.wikimedia.org/w/index.php?curid=44755052
The test is POSITIVE if the BD has generated a change (or %change) or the level of airways obstruction above clinically defined thresholds
Baseline assessment of airways obstruction*
Administration of a bronchodilator agent
(usually 400 mg of Salbutamol)
Post BD assessment
of airways obstruction
Wait (15-20 minutes)
PROCEDURE
1European Respiratory Society clinical practice guidelines for the diagnosis of asthma in children aged 5–16 years, European Respiratory Journal 2021 58: 2004173;
2European Respiratory Society Guidelines for the Diagnosis of Asthma in Adults, European Respiratory Journal 2022;

## Slide 51: Central and peripheral obstruction pre and post BD
NORMAL RANGES
(pred, ULN, LLN)
Rrs > ULN:
central airways obstruction
Xrs < LLN:
peripheral obstruction

## Slide 52: BDR testing using oscillometry
ERJ Open Res. 2021;7(4). 1

## Slide 53: BDR testing using oscillometry: ERS technical standard
The recommended thresholds for positive bronchodilator responses are:
Eur Respir J. 2020;55(2):1900753.
Param | Adults | Pediatric
ΔRrs% | -32% | -40%
ΔXrs% | +44% | +50%
Eur Respir J 2022; 59: 2102663

## Slide 54
Main clinical applications of oscillometry
Detection of airway obstruction
Bronchial reversibility and Bronchial provocation test
Detection of Tidal Expiratory Flow limitation (EFLt)
Other applications
Clinical applications of oscillometry

## Slide 55: Tidal EFL (EFLt): physiology
TIDAL EFL: the expiratory flows are so reduced that the lung cannot empty before the beginning of a new breath
“Expiratory flow limitation (EFL) is the pathophysiological hallmark of chronic obstructive pulmonary disease (COPD) and arises because of the dual effects of permanent parenchymal destruction (emphysema) and airway dysfunction, which in turn reflect the effects of small airway inflammation (mucosal oedema, airway remodelling and mucous impaction) and possibly increased cholinergic airway smooth muscle tone”
Physiology and consequences of lung hyperinflation in COPD.
D. E. O'Donnell, P. Laveneziana
European Respiratory Review 2006 15: 61-67;
Normal
COPD
Notes: 1- enfisema,.. Riduzione delle forze di richiamo eleastico
2- infiammazione, edema, smoth muscle.. La forza non si trasmette sul diametro interno la perdo sulle pareti

## Slide 56: Detection of EFL by FOT
EFL
DXrs
EFL  DXrs>2.8(4)
Inspiration
Expiration
Dellacà RL et al. Eur Respir J. 2004;23(2):232-240.
Notes: The measurements of within breath changes of input reactance been recently shown to be very effective for the detection of EFL in patients affected by COPD.
EFL is a fluidodinamic condition characterized by the development along the trechobronchial three of points in which flow is independent by the driving pressure. These points are called ccoke points.
The presence during expiration of the choke points prevents the stimulus from entering the alveoli, excluding them from the measurements. all that part of the lung that is downstream the flow limitnig sites.
Therefore during expiration reactance will be expression of the elastic propertyies of the whole respiratory system whiel during expiration reactance expresses the elastic properties of onle that parte of the lung that is upstream the flow-limiting sites
This woul d generate wide within-breath oscillation of the reactance and a paramenter called....
The presence of the choke points prevents the measurements of the whole system
The limiting case of the Venturi effect is when a fluid reaches the state of choked flow, where the fluid velocity approaches the local speed of sound. In choked flow the mass flow rate will not increase with a further decrease in the downstream pressure environment.
Dxrs therefore is a quantitative index of the amount of the airways that is chocked.

## Slide 57
Variability of within-breath reactance in COPD patients and its association with dyspnoea.
Aarli BB, Calverley PM a, Jensen RL, Eagan TML, Bakke PS, Hardie J a.
Eur Respir J. October 2014:1-10.
”ΔXrs was close to zero in healthy controls, with little within- and between-day variability. COPD patients tended to have higher DXrs values than the controls and more variation between measurements.
COPD patients with values exceeding those found in healthy individuals were prone to have higher mMRC scores at all GOLD grades.”
EFLt is an indicator of patient status that is relatively independent by the degree of obstruction measured with spirometry
Detection of tidal EFL using oscillometry

## Slide 58
Aarli et al - International Journal of COPD 2017:12 2179–2188
NO EFL, FEV1 >50%
EFL, FEV1 >50% !!!
“The present study suggests that evidence of EFLt, measured during tidal breathing in an effort-independent fashion, can identify a subgroup of COPD patients with worse clinical outcomes.”

## Slide 59
Main clinical applications of oscillometry
Detection of airway obstruction
Bronchial reversibility and Bronchial provocation test
Detection of Tidal Expiratory Flow limitation (EFLt)
Other applications
Mechanical ventilation
non-invasive ventilation
home-monitoring, longitudinal measurements
combination with slow vital capacity

## Slide 60: Intubate acute patients: evidences in newborns
Optimize ventilation parameters
- CDP - PEEP
Impact of treatments (Pre-post measures)
  Suction
  Surfactant
  Steroids
  Bronchodilators
  Posture
  Resuscitation at birth
  Etc. …
Longitudinal monitoring
Diagnosis - phenotyping
- Identifying deviation from “normality”
Prognosis

## Slide 61: Intubate acute patients: PEEP optimization
recruitment
XRS(5Hz) allows the definition of an optimal PEEP in agreement with CT data in an animal model of RDS
PEEP (cmH2O)
Xrs (cmH2O*s/L)
OPEEP
Dellacà et al 2011, Kostick et al 2011

## Slide 62: Automatic PEEP tailoring to abolish tidal expiratory flow limitation

## Slide 63: Longitudinal monitoring in acute patients
patients with moderate–severe COVID-19 ARDS.
Notes: Respiratory
Intensive Care Unit (RICU) COVID-19 of the Pulmonology Ward of the Azienda Sanitaria Universitaria Giuliano Isontina (Trieste, Italy) between

## Slide 64
Longitudinal home monitoring
Room-air oxygen saturation (SpO2) dropped to 82% on 6 November.
Blood testing (11 November) confirmed severe inflammation with an increase of D-dimer index (2500 ng·mL−1; normal values <500 ng·mL−1).
Home thoracic chest ultrasounds on 11 and 16 November demonstrated diffuse bilateral B-lines compatible with COVID-19 interstitial pneumonia.
The disease severity was “severe” per World Health Organization classification.
Notes: Impatto delle equazioni di riferimento

## Slide 65
“In this study, we have shown that the variability of FOT impedance measures, specifically those of Xinsp (SDXinsp and CVXinsp), is related to symptoms in COPD and may be used to detect changes prior to an AECOPD as early as 3 days before symptoms manifest.”.
“Feasibility of COPD home telemonitoring with FOT was high, consistent with previous studies [14, 36], with high adherence (>95%) and low drop-out rates (<7%)”
Beginning of symptoms
Beginning FOT changes
Exacerbation

## Slide 66: Longitudinal measurements
Monitoring the Temporal Changes of Respiratory Resistance: A Novel Test for the Management of Asthma. American journal of respiratory and critical care medicine. 2012

## Slide 67
Picture from: https://derangedphysiology.com/
Oscillomety + slow vital capacity: Detection of Closing volume – Specific airway conductance

## Slide 68
Oscillomety + slow vital capacity: Detection of Closing volume

## Slide 69: Oscillomety + slow vital capacity Specific airway conductance (sGrs)
“In our cohort, the sGrs outperformed the spirometric and the FOT indices Rrs5 and Xrs5 regarding their relationship with N2‐MBW outcomes.
Notably, the relationship between LCI and sGrs remained unaffected by lung hyperinflation, as opposed to that between LCI and FEF25‐ 75, Rrs5, or Xrs5, which was significantly reduced at higher RV/TLC values (Figure 4).”
Pediatric Pulmonology

## Slide 70
Thank you!
chiara.veneroni@polimi.it

## Slide 72: How should I integrate oscillometry in my clinical routine?
Integration of knowledge of physiological determinants of test results into a functional classification of the identified impairments.
1
2
3
Integration of the identified patterns with other clinical data to inform differential diagnosis and guide therapy.” ERS/ATS technical standard on interpretive strategies for routine lung function tests. 2021
“Consideration of the measurement error of the test, as well as the inherent biological variability of measurements both between individuals and between repeated measurements in the same individual.
Integration into clinical practice of oscillometry testing should be similar to how one would use other lung function tests:
Are Rrs parameters the ONLY abnormal ones?
(indicative of possible reduction of the caliber of the central airways)
Are Xrs parameters the ONLY abnormal ones?
(indicative of possible small, distal airways dysfunction/airway closure/lung
parenchymal diseases)
Are BOTH Rrs and Xrs abnormal?
(indicative of possible mixed reduction of the caliber of the central airways and small,
distal airways dysfunction)
Is EFLt present (i.e. ΔXrs ≥ 2.8)?
(indicative of the presence of dynamic hyperinflation
Is Ve of the session ≤ 18 L/min?
Is CoV% of Rtot of the session ≤ 10% (adults) or ≤ 15% (children)?
Oscillometry results: